Applied BioSystems
MDS Sciex

Mascot Script
User Manual

# 1 Introduction

This document describes a script for sending either the active Analyst spectrum or all product spectra contained in the active sample (or all samples in the active data file) to the Mascot protein search engine. The script was co-developed with Matrix Science Limited – the creators or Mascot.

When only sending the active spectrum, the script can work with either MS or MS/MS data; in the former case a peptide mass fingerprint search is conducted. When sending all spectra, the script works with data acquired using two distinct types of acquisition method: either a multiple period / multiple experiment method containing any number of product experiments or an IDA method. In the former case the script calculates one spectrum for each experiment by averaging all spectra acquired for the experiment. In the latter case the script uses all dependent product spectra, combining adjacent spectra with the same parent mass and charge state.

# 2 Operation

If you are planning to send just one spectrum to Mascot, ensure that a centroided spectrum is active in Analyst before running the script. If you are planning to send all spectra contained in a sample, the TIC for the sample should be active. The following figure shows the resulting dialog.

Mascot Search Dialog

The dialog contains the following items:

Search Pane – There are four different options available:

The current spectrum – If the active Analyst pane is a centroided spectrum (either MS or MS/MS) this option is used to search this spectrum on its own. In order for this option to be available the spectrum must have been centroided using the ‘Centroid’ command from Analyst’s ‘Explore’ menu.

All MS-MS spectra for current sample – If the active Analyst pane contains MS/MS data this option is used to perform a single search using all spectra for the sample.

All MS-MS spectra from selected region in TIC – If the active Analyst pane contains MS/MS data and also contains a selected region, this option is used to perform a single search using only the spectra from this region. This can be particularly useful to speed the generation of the search input if only a portion of the run is known to contain useful data.

All MS-MS spectra from all samples for current file – If the active Analyst pane contains MS/MS data and the associated data file contains more than one sample, this option is used to perform a single search using all MS/MS spectra from all samples in the data file.

File, Sample, Periods, Scan type, IDA data – These read-only fields show information about the active Analyst sample.

Cancel – Click this button to dismiss the dialog. If a search operation is in progress it will be aborted before the dialog is dismissed.

Options – Click this button to set the various search options. See below for more details.

Search – Click this button to open the Mascot search form and populate it with the appropriate information. Note that if you are searching all product spectra contained in the sample this may take some time (a progress bar will appear). Once the web form appears, you will need to manually click its ‘Start Search’ button.

Mascot Search Options

The figure above shows the ‘Options’ dialog; it contains the following items:

Mascot parameters

Set default parameters – Clicking this button displays the ‘Mascot search form defaults’ web page. You can edit the various defaults so that you do not need to reset them manually every time before submitting a search. After changing the parameters, click the ‘Save defaults as cookie’ button and close the web page.

Set search file location – Click this button to set the location of the text file generated by the script and used as input to Mascot. For details, see below.

Web Server Location – Specify where the search engine is located. If you have installed Mascot onto your local computer, you may choose the ‘Mascot on this computer’ option (otherwise this option is not available). Otherwise specify whether to use the Matrix Science public web site or a local Mascot or Protein Prospector intranet site (and specify the actual address).

Try to determine charge state from survey scan – If this option is unchecked, the script will always use the default precursor charge states as-is. Otherwise for QStar data it will attempt to determine the charge state by examining the isotope spacing in the survey spectrum; for Analyst QT data it will use the charge state determined by the MS Acquisition Engine (which is saved to the data file). If this option is checked but the charge state determination fails, the default charge states are used.

Default precursor charge states

When searching only the current spectrum, the precursor charge state should be set manually using the resulting Mascot web form and the values here are not available. When searching product spectra for an IDA run, the charge state is automatically determined by the script – the values specified here are not used unless the charge state could not be automatically determined or the ‘try to determine charge state from survey scan’ option was not selected. When searching product spectra for a multiple period / multiple experiment LC/MS run, the specified charge states will always be considered for each product spectrum.

Discard ions with charge of 5+ or higher – Some versions of the Mascot search engine cannot accept ions with a charge state of five or higher. At the time of this writing the version at the Matrix Science public web site can accept higher charge states, but this version does not (yet) appear to be widespread on local intranets. If you are using one of these older versions you should make sure that this option is checked. If you do not, you will see a warning after submitting the Mascot search for each ion with a charge state of five or higher.

MS/MS averaging of IDA dependents – These fields allow you to specify parameters which pertain to the calculation of the product ion spectra for an IDA run.

Precursor mass tolerance for grouping – This field is used to potentially combine adjacent product spectra into a single spectrum. If two spectra have precursors with the same charge state and the same MW – within this tolerance – they will be combined.

Max. number cycles between groups – Spectra are not combined if the number of cycles between spectra with the same m/z and charge state is greater than the specified value. The idea is that spectra with sufficiently different retention time should not be combined.

Min. num cycles per group – Specifies the minimum number of spectra which need to be combined in order for the result to be kept. Note: If you have used the dynamic exclusion IDA option, you must set this value to ‘1’.

MS/MS data processing – These fields allow you to specify how the product spectrum are processed. In the case of an IDA run, this processing happens after any adjacent spectra are combined.

##### Threshold peaks with intensity < ‘x’ –  In general it is recommended to keep all peaks found by the centroid algorithm and to use a value of zero here. However if desired you can enter a larger value to remove small peaks from the MS/MS spectra. If the ‘(absolute)’ setting is selected the threshold is in actual counts (for quadrupole data) or counts/sec (for TOF data), otherwise it is expressed as a percentage of the base peak for the spectrum.

Centroid all MS/MS data – Specify whether or not to centroid the MS/MS spectra before sending them to Mascot for searching. It is highly recommended to enable the centroid option.

De-isotope MS/MS data – If the centroid option is used, specify whether or not isotope peaks should be removed from the MS/MS spectra before sending them to Mascot. It is recommended to enable this option.

Reject spectra if less than ‘n’ peaks – If a spectrum contains unreasonably few peaks after combining adjacent spectra (if used) and centroiding (if used) it can be eliminated.

Search File Location Options

The figure above shows the dialog which is presented when you click the ‘Set search file location’ button in the options dialog. It contains the following items which allow you to specify where the text file used as input to Mascot will be located:

In the Windows ‘Temp’ folder with a random filename – If this option is selected the file is placed into the Windows temporary folder; the filename will be random, but unique.

Always to the following file – This option will always write to a specific file, overwriting the file for every search.

To the following directory with a random filename – This option will create the file at the directory which you specify with a random (but unique) filename.

Prompt me each time for the location – This option will cause a standard ‘Save As’ dialog to be presented every time a search is performed asking you for the location of the file.